晶准/Jing Zhun		 数显指示表使用说明书/User’s Manual of Digital Indicator

一、性能指标/Performance index

1、公英制转换/ Inch/Metric measurement

2、绝对/相对状态转换/Absolute Measuring/Relative measuring

3、外部清零/Zero Setting

4、正/负方向转换/ Switch between +/-

5、自动/手动关机/Auto/manual Power Off

6、预设实际位置数值/Preset a Value

7、跳动量高速检测/Measure run out of a high speed work piece

8、数据输出协议兼容三丰量表/The protocol of data output is compatible with Mitutoyo indicator

二、产品外观/Product appearance

| - 检测头 - Anvil - 夹持套 - Stem - 电池盒 - Battery compartment - 后盖 - Rear cover 5、LCD 6、数据输出口 Data output port - C键 - Clear button - R健 - Run button - in/mm键 - in/mm button - ABS/PRESET键 - ABS/PRESET button | - Z/ON键 - Z/ON button - 公制单位标志 - Display Metric - 英制单位标志 Display Inch - 预设值标志 - Display Preset - 跳动量测量标志 - Display Graduation - 绝对状态标志 - Display ABS mode - 低电压标志 - Display Low voltage - 最大值标志 - Display Max. Value - 最小值标志 - Display Min. Value |
| --- | --- |

三、技术规格/Technical specification

Battery：3V锂电池CR2032

Power supply: 3V lithium battery CR2032

工作温度：0~40°

Working temperature: 0~40° DEGREE CENTIGRADE

数据输出：RS232				储运温度：-10°~60°

Data output: RS232     Storage and transport temperature: -10°~60° DEGREE CENTIGRADE

四、注意事项/Precautions

1、使用前请先用不含油的柔软干布对表壳、测量杆等地方进行擦拭。如果有严重的污垢，请用含有中性清洁剂的抹布进行擦拭。

Clean the indicator housing, the stem and the testing rod,etc with soft oil-free cloth.

If it’s very dirty, put some neutral cleaner on the cloth.

- 请使用夹持孔规格为8H6mm的夹具来夹持量表。

The hole that holds the indicator stem should be 8H6mm.

- 使用完后请同样清洁表壳和测杆等地方。

Please also clean the indicator housing, the stem and the testing rod,etc after using.

- 请不要自行拆装量表，以免造成严重后果。

- Please do not disassemble the indicator by yourself to avoid serious consequences.

- LCD显示变暗，数字不清晰或无法开机时请更换电池。

- Replace the battery when LCD turns dim, the figures are not clear or the indicator

- cannot be turned on.

五、功能介绍/Functions

1、开机和关机/Turn on and Turn off

a）装上电池后直接开机。

Turn on the indicator directly after installing the battery.

b）短按键可以实现开机。

Shortly press to turn on.

c）长按键实现关机。

Long press  to turn off

.

2、清零操作/Zero setting

a）ABS测量模式下（屏幕显示ABS）短按键实现清零，或者回到设定好的初始值。

In ABS Mode (ABS is shown on the screen), press for a short duration, the screen will show “0”, or the zero point that you have set.

b）从ABS绝对测量模式转到相对测量模式时，自动清零。

When the indicator is turned from ABS mode to Relative mode, the screen will show zero and you may achieve relative measuring.

c）在相对测量模式状态下不能实现清零操作。

When the indicator is in Relative mode, zero cannot be reset.

3、公英制转换/Inch/metric

短按键可进行公英制单位的转换。显示mm符号为公制单位；显示in为英制单位。

Shortly press to switch between Inch/metric. If mm is displayed, the value is in metric. If inch is displayed, the value is in inch.

公制单位

Metric

英制单位

Inch

4、绝对测量和相对测量模式转换/Switch between Absolute Measuring and Relative Measuring

在基本状态下短按键可以转换相对测量和绝对测量状态，显示“ABS”符号是为绝对测量状态。转换后在相对状态下显示为相对零点，在绝对状态下则显示为测杆当前的绝对数值（例如图中的8. 88mm）。

Shortly press under the basic status to switch between Relative mode and ABS mode, and the "ABS" indicator is for absolute measuring mode.After switching, it shows a relative zero point under relative measuring mode, and it shows a current absolute value of the measuring rod under the absolute measuring mode (for example, 8.88mm in the legend).

绝对测量状态

Absolute Measuring

相对测量状态

Relative Measuring

5、正/负方向转换/Switch between +/-

长按键可使量表数据变化与测杆移动方向的相对关发生改变。

Long press button to make the relative change of the indicator data and the movement direction of the measuring rod.

长按键负号闪动，等负号稳定消失后，推动传感器数据由零向负数变化：

Long press button until “-” code blinks, and after “-” stops blinking and disappears, push the Anvil, the value will be minus:

推动测杆/Push the Anvil

长按键正号闪动，等正号稳定消失后，推动传感器数据由零向正数变化：

Long press button until “+” code blinks, and after “+” stops blinking and disappears, push the Anvil, the value will be plus:

推动测杆/Push the Anvil

6、预设实际位置数值/Preset a value

在公制的绝对状态下，长按键进入预设值状态，最右一位数字闪动，闪动的数值就是可以改写的数值。短按键数据由0到9变化，短按键数据向左一位并且闪动，依次类推设定最左一位数字正负号。设定好后，长按键退出并保存预设值。退出设置状态后，量表将自动进入绝对测量状态，此后绝对零位显示数值为预设的数值。在设置过程中长按键可以把已设的数值快速清除，并退出设置状态。（注意，预设值不能超过99!）

When the indicator is in “ABS” and “metric” mode, long pressbutton to enter the presetting mode, and the rightmost number blinking. The number that is blinking is the one that can be changed. Shortly pressbutton, then the number changes from 0 to 9, shortly press button, then the blinking number will move one position left. In this way all the numbers and +/- can be set. After setting, long press button to save the preset value and quit presetting mode. After quitting, the indicator will automatically enter the absolute measuring mode, and now the real absolute Zero point is with the value preset. Long press  button to quickly zero the preset value during the setting and quit.(Note: the preset value cannot be higher than 99!)

7、跳动量高速检测/Measure run out of a high speed work piece

量表全程自动记忆存储最大最小值，短按键可清除历史记忆，液晶显示〃-CCC-"，重新跟踪跳动。短按键查看量，显示RAN字符及符号查看最大值。再短按键时显示符号，此时显示为最小值。然后短按键显示符号，此时显示最大最小值差。短按键返回普通测量界面。

The indicator may record max. and min. value automatically of the whole process, shortly press button to erase the memory, the LCD will display “-CCC-", and will record again. Shortly pressbutton to check the run-out.When RAN and  are displayed, the value shown on LCD is the max. Run-out. Shortly press button again,  will be displayed and the value on LCD is the min. Run-out. Then shortly pressbutton again.  will be displayed. The value shown on LCD is the difference between maximum value and minimum value. Shortly press button to return to the normal measuring interface.

注：跳动量检测可停留在三种状态中的任何一种，做独立检测，且在各状态下短按键能清除记忆。

Note: The run-out detection can stay in any of the three modes for independent detection, say min. Run-out or max. Run-out or max. & min. Run-out. And shortly press  button to erase the memory in any mode.

8、手动/自动关机切换/Switch between Manual/Auto Power Off

关机状态下，长按键不松手，此时按开机键，可进入手动/自动关机设置状态，“H”表示Hand手动关机，长按键关机；“A”表示Auto自动关机，在测杆没有移动以及无按键操作情况下，4小时后自动关机。短按键保存设置并退出设置状态。

When the indicator is turned off, long press  button without releasing, and then press the power button to enter the Manual/Auto Power off Setting mode. “H” means Hand (manual) power off, the indicator is powered off when   is pressed for a long duration. "A" means Auto for automatic power off, the indicator will automatically turn off in 4 hours when  If testing rod stays still and no button is pressed. Shortly press button to save the setting and quit.